## Slide 1: An alternative sensor-based method for glucose monitoring in children and young people with diabetes
Edge J, et al. Arch Dis Child. 2017;102(6):543–549

## Slide 2: Introduction
While safe and efficacious, CGM is underused in paediatric patients with diabetes, despite >75% not meeting ISPAD/ADA guidelines for glycaemic control
Among the barriers to CGM use are:
  weekly sensor insertion and multiple SMBG tests for calibration
  discomfort/pain at the insertion site and
  loss of data due to poor sensor–receiver connectivity
This study evaluated the accuracy, safety and patient acceptability of the FreeStyle Libre™ Flash Glucose Monitoring System in children/young people with diabetes
ADA, American Diabetes Association; CGM, continuous glucose monitoring; ISPAD, International Society for Paediatric and Adolescent Diabetes; SMBG, self-monitoring of blood glucose Edge J, et al. Arch Dis Child. 2017;102(6):543–54

## Slide 3: FreeStyle Libre™ Flash Glucose Monitoring System
Factory-calibrated sensor
Worn for 14 days without needing calibration
Automatic storage of glucose data every 15 minutes
Real-time glucose scanning (with 8-hour trend), can also be conducted as often as required
Glucose reports can be reviewed by patients at home or in the clinic with healthcare professionals
Edge J, et al. Arch Dis Child. 2017;102(6):543–54

## Slide 4: Study design
A prospective, single-arm study, conducted in 9 UK centres (ClinicalTrials.gov identifier: NCT02388815)
AEs, adverse events; BG, blood glucose; CGM, continuous glucose monitoring; CSII, continuous subcutaneous insulin infusion; MDI, multiple daily injections Edge J, et al. Arch Dis Child. 2017;102(6):543–54
Flash sensor-based glucose monitoring
Participants performed four capillary BG tests daily (pre-existing diabetes self-management plan maintained)
Children and adolescents (4–17 years old) with type 1 or 2 diabetes; receiving MDI insulin or CSII; BG testing ≥2/day; not currently using a CGM device
Visit 1 Physical examination, device training, sensor
insertion, AEs, sensor application questionnaire
Visit 2 (days 5–8) Safety monitoring check up, AEs, data upload
Visit 3 (days 12–15) AEs, device unmasking, sensors removed, site inspection, data upload, user questionnaire
14-day study phase
Screening and enrolment

## Slide 5: Outcomes
The primary outcome was sensor accuracy compared with capillary BG measurements
Secondary outcomes included:
  Glycaemic variability
  User acceptability
  Adverse events
BG, blood glucose Edge J, et al. Arch Dis Child. 2017;102(6):543–54

## Slide 6: Participants
Chart: 
| Category 1
Series 1 | 149
Chart: 
| Category 1
Series 1 | 139
Series 2 | 62
BG, blood glucose Edge J, et al. Arch Dis Child. 2017;102(6):543–54
Chart: 
| Category 1
Series 1 | 149
A total of 89 participants were enrolled
Chart: 
| Category 1
Series 1 | 139
Series 2 | 62
Paired sensor–capillary BG results were not available for 2 participants
Withdrawal prior to obtaining sensor data (n=1)
Protocol deviation (n=1)
89 enrolled
87 included
3 excluded
2 excluded
87 evaluable participants included in accuracy analyses
89 participants included in safety analyses
88 participants included in glycaemic variability analyses
88 included
89 included

## Slide 7: Baseline characteristics
CSII, continuous subcutaneous insulin; HbA1c, glycated haemoglobin; MDI, multiple daily injections; SD, standard deviation Edge J, et al. Arch Dis Child. 2017;102(6):543–54
Characteristic | n=89
Mean age, years | 10.4
Age categories, n (%) | 
4–7 years | 24 (27.0)
8–12 years | 39 (43.8)
13–17 years | 26 (29.2)
Male, n (%) | 45 (50.6)
Type 1 diabetes, n (%) | 89 (100.0)
Mean ± SD duration of diabetes, years | 4.0 ± 2.8
Mean ± SD HbA1c, % (mmol/mol) | 7.6 ± 1.1 (60 ± 12)
Insulin regimen, n (%) | 
CSII | 50 (56.2)
MDI | 39 (43.8)

## Slide 8: Sensor accuracy
Sensor results were in good agreement with capillary BG results
ARD, absolute relative difference; BG, blood glucose; CEG, consensus error grid; MAD, mean absolute difference; MARD, mean absolute relative difference; MRD, mean relative difference Edge J, et al. Arch Dis Child. 2017;102(6):543–54
 | Accuracy
Overall sensor accuracy (n=5493*) | 
MARD | 13.9%
Median ARD | 10.4%
MRD | 1.7%
%within1.1/20 | 82.0%
% within CEG Zone A | 83.8%
% within CEG Zone B | 99.4%
MAD at glucose concentrations <5.55 mmol/L (n=1468*) | 0.75 mmol/L
MARD at glucose concentrations 5.55–10.0 mmol/L (n=2090*) | 13.5%
MARD at glucose concentrations >10.0 mmol/L (n=1935*) | 10.6%
*Paired sensor–capillary BG results. %within1.1/20, proportion of results within ± 1.1 mmol/L of the BG value for glucose <5.55 mmol/L and within ± 20% of the BG value for glucose ≥5.55 mmol/L | 

## Slide 9: Effects of baseline variables on sensor accuracy
There were no statistically significant differences in clinical accuracy detected for:
  Patient age (p=0.133)
  Patient sex (p=0.951)
  Body weight (p=0.284)
  Method of insulin administration (p=0.640)
  Sensor lot (p=0.135)
  Daytime vs night-time use (p=0.909)
Edge J, et al. Arch Dis Child. 2017;102(6):543–54

## Slide 10: Glycaemic variability
*06:00 to 22:00 †22:00 to 06:00 CSII, continuous subcutaneous insulin infusion; MDI, multiple daily injection Edge J, et al. Arch Dis Child. 2017;102(6):543–54
Patients were in euglycaemia for ~50% of the day
Chart: Mean time in range (hours)
| Glucose <3 mmol/L | Glucose <3.9 mmol/L | Glucose 3.9–10 mmol/L | Glucose >10 mmol/L | Glucose >13.3 mmol/L
Overall (24 hours, N=88) | 0.9 | 2.2 | 12.1 | 9.6 | 5.2
Daytime (N=88) | 0.4 | 1.2 | 8.1 | 6.7 | 3.8
Night-time (N=87) | 0.5 | 1 | 4.1 | 2.9 | 1.5
CSII (N=50) | 0.7 | 2 | 12.1 | 9.9 | 5.3
MDI (N=38) | 1.3 | 2.6 | 12.2 | 9.2 | 5.1
Time in hypoglycaemia
Time in hyperglycaemia
*
†
Time in euglycaemia

## Slide 11: User acceptability
The system was rated favourably* by most respondents, and was perceived similarly by boys and girls
*Responses rated on a scale of 0 (strongly agree) to 4 (strongly disagree), with favourable acceptability defined as a score of 0 or 1; SMBG, self-monitoring of blood glucose Edge J, et al. Arch Dis Child. 2017;102(6):543–54
Question topic | % patients scoring questions favourably*
Sensor application | 84.3–92.1
Sensor wear and use | 87.2–100
Comparison with SMBG | 85.4–97.5
The device | 68.3–96.3

## Slide 12: Adverse events
One SAE unrelated to the study or device was reported (pain and lack of feeling in leg)
5 device-related AEs (mostly mild) occurred:
  allergic reaction, blister and pink mark/scabbing in 1 patient each
  abrasion on sensor removal in 2 patients
  all were resolved at study completion
Anticipated sensor application-associated AEs included:
  moderate erythema (11.6% of site insertions)
  mild erythema (13.6%)
  pain (4.1%)
  mild bleeding, bruising, itching and oedema (<3% each)
AEs, adverse events; SAE serious adverse event Edge J, et al. Arch Dis Child. 2017;102(6):543–54

## Slide 13: Conclusions
The FreeStyle Libre™ Flash Glucose Monitoring System was accurate and safe, and had favourable acceptability in paediatric patients (4–17 years old) with type 1 diabetes
Accuracy was unaffected by patient characteristics, making it suitable for a broad range of children and young people with diabetes
Edge J, et al. Arch Dis Child. 2017;102(6):543–54